| Testing Requirements # Parameter | # Example | # Instrument |
| --- | --- | --- |
| # Optical power | # Source output, receiver signal level | # Power meter |
| # Attenuation or loss | # Fibers, cables, connectors | # Power meter and source, or Optical Loss Test Set (OLTS) |
| # Back reflection or Optical Return Loss (ORL) |  | # OTDR or OCWR |
| # Source wavelength |  | # Spectrum analyzer |
| # Backscatter | # Loss, length, fault location | # OTDR |
| # Fault location |  | # OTDR, VFL |
| # Bandwidth/dispersion |  | # Bandwidth tester |

# Power Meters

## The power meter by itself can be use to measure source power

## With a source, it can measure the loss of a cable plant, called insertion loss

## Most power measurements are in the range +10 dBm to -40 dBm

### Analog CATV (cable TV) or DWDM (Dense Wavelength Division Multiplexing) systems can have power up to +30 dBm (1 watt)

# Wavelengths

## Power meters are calibrated at three standard wavelengths

### 850 nm, 1300 nm, 1550 nm

## Typical measurement uncertainty is 5% (0.2 dB)

Sources

Images from lanshack.com

## Sources are either LED or laser

### 665 nm for plastic optical fiber

### 850 nm or 1300 nm for multimode

### 1310 nm or 1550 nm for singlemode

## Test your system with a source similar to the one that will be actually used to send data

# Optical Loss Test Set

## Power meter and source in a single unit

### Normally used in pairs

### Automated, more complex and expensive than the combination of a source and a power meter

#### Rare in field testing

##### Image from aflfiber.com

# OTDR (Optical Time-Domain Reflectometer)

## Image from exfo.com

# OTDR Uses

## Measure loss

## Locate breaks, splices, and connectors

## Produces graphic display of fiber status

### Can be stored for documentation and later reference

## Cable can be measured from one end

# Backscatter

## A small amount of light is scattered back to the source from the fiber itself

## Splices or connector pairs cause a larger reflection of light back to the source

### Figure from techoptics.com (link Ch 17a)

Dead

zone

# OTDR Display

# OTDR Accuracy

## OTDR can give false loss values when coupling different fibers together

### Splices can even show more light on the other side “gainer”

### This is an illusion caused by increased scattering on the other side

### Splice loss uncertainty up to 0.8 dB

# Types of OTDR

## Full-size

### Complex, powerful, expensive

## Mini-OTDR

### Fewer features

## Fault Finder

### Simplified, shows distance to a fault

#### Links Ch 17c, d, e

# Visual Cable Tracers and Visual Fault Locators

## Cable tracer is just a flashlight

## VFL uses an LED or Laser source to get more light into the fiber

### Useful to test a fiber for continuity

### To check to make sure the correct fiber is connected

### With bright sources, you can find the break by looking for light shining through the jacket

#### Visible light only goes 3-5 km through fiber

##### Images from links Ch 17 e & f

# Fiber Identifiers

## Bends the fiber to detect the light

## Can be used on live fiber without interrupting service

## Can detect a special modulated tone sent down a fiber

### Image from tecratools.com (link Ch 17d)

Cable to be Tested

# Optical Continuous Wave Reflectometer (OCWR)

## Measures optical return loss (reflectance) of connectors

## Inaccurate on installed systems because it includes backscatter and all sources of reflectance

### See link Ch 17h

# Microscope

## Used to inspect fibers and connectors

### Particularly during epoxy-polish process

#### Image from link Ch 17g

# Talkset

## Telephone calls over unused fibers

## Rarely needed now that we have cellphones

### See link Ch 17i

# Attenuators

## Simulates the loss of a long fiber run

## Variable attenuators allow testing a network to see how much loss it can withstand

## Can use a gap, bending, or inserting optical filters

### Image from link Ch 17j

# Reference Cables

## Test cables are needed to connect the cables to be tested to the test instruments

## Must have correct connectors, be clean, and high-quality (low loss)

## Use high-quality mating adapters

### Ceramic or metal – not plastic

### Singlemode rated are most accurate

# Optical Power Levels

| Detectors are Silicon, Germanium, or Indium-Gallium-Arsenide semiconductors # Network Type | # Wavelength | # Power Range (dBm) |
| --- | --- | --- |
| # Telecom | # 1330, 1550 | # +3 to -45 |
| # Telecom DWDM | # 1550 | # +20 to -30 |
| # Data | # 665, 790, 850, 1300 | # -10 to -30 |
| # CATV | # 1300, 1550 | # +10 to -6 |

# Calibrations

## NIST is a standards laboratory

### Offers power calibration services at 850, 1300, and 1550 nm wavelengths

### Instruments should be returned to the manufacturer for calibration annually

# Uncertainties

## Absolute power: 5% or 0.2 dB

## Insertion loss: 0.5 dB or more

## OTDR: up to several dB

## Optical return loss: 1 dB or more

### Although meters show a reading with hundredths of a decibel, they don’t mean anything

#### A 2.13 dB loss might well re-measure as 2.54 dB

# Optical Fiber Testing

## Before installation

### Test continuity with cable tracer or VFL

## Measure attenuation with cutback method

### Cut off known length, measure power increase

# Sources for Loss Measurements

## Most multimode systems use LED sources

### High-speed multimode often uses VCSELs (1 Gbps and higher)

### See link Ch 17k

## Singlemode systems use laser sources

## Test with the source you will really use

### BUT Argilent says you should test all Multimode with LEDs (link Ch 17l)

# Modal Effects in Multimode Fiber

## Mode scramblers mix modes to equalize power in all modes

### Can be made with a section of step-index fiber

## Mode filters remove higher-order modes to reach equilibrium modal distribution

### Can be made with a mandrel wrap

## Singlemode fibers shorter than 10 meters may have extra modes

### Use a launch cord to avoid that problem

# OTDR Pulse Width

## Longer pulses can see further down the cable because they have more light

## But they have less accuracy finding locations

### From link Ch 17a

# OTDR Uncertainties

## Dead zone

### Nothing can be measured for the first 100 meters or so

## Distance Resolution

### Two events too close together cannot be resolved

### Especially with long pulses

# OTDR Distance Errors

## Speed of light in fiber

### May not be exactly what the OTDR expects, distorting distances

## Slack in fiber

### OTDR measures length along the fiber, which is usually 1% - 2% longer than the length along the cable

# OTDR Loss Errors

## Joining two fibers with different backscatter coefficients will cause:

### Too high a loss when measured in one direction

### Too low a loss in the other direction

## For accurate loss measurements, measure from both ends and average the results

# OTDR Ghosts

## Secondary reflection appears at double the real cable length

## Using index-matching gel will eliminate ghosts

# Dispersion

## Multimode fibers suffer from modal dispersion

## All fibers suffer from chromatic dispersion

### Because different wavelengths travel at different speeds, and no source is completely monochromatic

## In very long singlemode networks, polarization mode dispersion also matters

# Bandwidth Testers

## There is a new unit available to test bandwidth in the field, but it is not commonly done yet (link Ch 17 k)

# Connector Insertion Loss Test

## This test gives the typical loss of a connector type

# Modal Distribution

## The insertion loss test

### FOTP-34 by the TIA

## Three options of modal distribution

### EMD or steady state

#### After a mandrel wrap

### Fully filled

#### After a mode scrambler

### Any other specified conditions

# Microscopes

## Used to inspect the ends of polished connectors

## Helpful to view the connector at an angle while lighting it from the side

## Only defects over the core really matter

# Optical Return Loss in Connectors

## A pair of glass-air interfaces for nonphysical contact connectors without index-matching gel

### 4% reflectance – loss of 0.3 dB

## PC connectors can have a reflectance of 1% or an ORL of 20 dB

### Much less with the best PC connectors – 40 to 60 dB

## Reflectance can be a problem in high bitrate singlemode systems

# Basic Cable Loss Test

## Test FOTP-171

### Measure power through launch cable

### Then add cable to test

#### This tests only one connector – turn the cable around to test the other end

# Double-Ended Loss Test

## Uses both a launch and receive cable

# Single-Cable Reference

## Refer to this condition:

## Test this way

### EIA/TIA 568 requires this

### See link Ch 17m

# Why Use Single-Cable Reference?

## It gives highest loss and lowest uncertainty

# Choosing a Launch Cable for Testing

## Choose cables with low loss

### It is not necessary to get connectors and fiber with tighter specifications

## Handle the launch cables carefully

## Inspect them with a microscope

## Keep them clean

### Use splice bushings with metal or ceramic alignment sleeves

# Mismatched Fibers

## Coupling a smaller fiber to a larger one causes only a small loss (0.3 dB or so)

## Connecting large fiber to small fiber causes a large loss

### Both because of diameter and numerical aperture

# Testing the Installed Cable Plant

## Can use one-cable reference, or two-cable, or three-cable, but the type of reference must be documented

# Wavelengths

## Usually test multimode at both 850 and 1300 nm with LED sources

## Singlemode test is usually at 1300 nm only

### 1550 nm is sometimes required also

### For long-distance, and to show that WDM can be performed later

### Also shows microbends – 1550 test is much more sensitive to bending loss

# Optical Splitter

## Splits light signal from one fiber into two fibers

### Figures from tpub.com (link Ch 17n)

# Couplers Can Split or Combine

## You can also split one to M, or combine M to 1

# M to N Coupler

# Making Couplers

# Wavelength Division Multiplexers

## Light entering from the left containing two wavelengths is separated into the two fibers on the right

## Combining the two signals is also possible

## Requires special equipment and techniques to test

### Image from link Ch 17o

# Fiber Optic Amplifiers

## Boosts signal without converting it to electricity

## Complicated to test, require special equipment

### Image from link Ch 17p

# Fiber Optic Switch

## See links Ch 17q and 17r

# Fiber Optic Datalinks

## The diagram shows a single link

## Most networks will be bidirectional (full duplex) with two links working in opposite directions

# Bit Error Rate

## The receiver power must be within the operating range

### Too little power leads to high bit error rates (wrong data at receiver)

### Too much power saturates the detector and also leads to high bit error rates

#### Use an attenuator in this case

# What Goes Wrong?

## Often the two fibers are connected backwards – check them with a visual tracer

## Check receiver power level

## Check plant loss with double-ended method

# Don’t Use an OTDR to Measure Plant Loss

## OTDR does not see the loss of the end connectors

## Its power source is not the same as normal LAN power sources

## OTDR measurements are affected by backscatter coefficient which may not be the same for all the cables in a network

# Back Reflection

## Back reflection can cause networks to fail even though the loss is low

## Power meter can’t measure reflection

### Use an OTDR or OCWR

### Cure it by splicing in low-reflection patch cords to replace high-reflectance connectors

### Angled PC connectors are designed to minimize reflectance for this reason (not mentioned in textbook)

# Reliability

## Once installed, the fiber optics should work for a long time

## People break the cable by accident

### Mark where cables are buried

### Bury a marker tape above the cable

### Use orange or yellow jacket cable indoors

### A broken cable just behind a connector in a patch panel is hard to find

# Source Failure

## LED in laser transmitter drops in power as it ages

## Laser sources are feedback-stabilized so they remain constant in power till they fail

Last modified 5-6-06