Optical Time Domain Reflectometer (OTDR)

The OTDR sends a brief pulse of infrared light into the fiber, and measures the back-scattered light that returns from the cable.  This gives a graphical display of the fiber length, showing connectors, splices, and the end of the cable.  It is especially helpful in locating a break in a long fiber link.

Goal:

The goal is to get a good trace of the CCSF fiber loop, which is about 60 km long.

## Starting the EXFO OTDR

- Turn it on with the            button on the lower right of the front panel

- The screen dims a lot, just press the            button several times to make it brighten again.

- The starting screen shows the module we have installed:

### FTB-7300D-23B-EI-EIUL	This is the OTDR.

## Collecting a Trace

- On the bottom of the screen, there are three buttons:

Advanced OTDR	Automatic OTDR	Create Ref./Template

- Touch the Automatic OTDR button on the screen with the blue stylus.  Don't use the tip of a pen because that will mark up the screen.

- Touch the square Start button in the lower right of the screen.

- An old trace is shown.  Ignore it and connect the fiber to the CCSF network at the wall-mounted patch panel.

- In the Wavelength section, notice that both 1310 nm and 1550 nm  are selected with X signs.  That’s OK – we will use both wavelengths.

- Make sure the three dials: Distance (km), Pulse, and Time (s) are all set to Auto.

- Press the green Start button in the top right.  If a box appears saying “Data Unsaved”, click No.  While the trace runs.  You will see a yellow warning box saying Warning!  Laser is ON in the lower right of the screen.  It will take 30 seconds for each wavelength.

- When the trace is collected, you will see a trace on the screen.  Compare it to the figures on the next page.

- A good trace will show a spike at 0 km and a gradual slope going out to 66 km or so, with spikes in between at each patch panel.

- A bad trace will just fade to noise within 0.5 km, indicating that the launch cable is incorrectly connected.

Last modified 1-10-07

Example OTDR Traces

Start of cable

Connection Points

End of cable

Good Trace of

66 km CCSF Loop

Bad Trace:

Launch Cable Improperly Attached